## Slide 1
Engineering Science Thesis
Sleep staging AI model design and ASIC implementation
Tristan Robitaille (EngSci 2T3 + PEY) Supervisor: Prof. Xilin Liu (X-Lab)
Fall 2023 – Winter 2024

## Slide 2
Fall 2023 – Winter 2024
Outline
Background
Literature Review
Model Description
Accelerator Design
Results
Analysis & Future Work
Conclusion

## Slide 3
Background
Fall 2023 – Winter 2024
1
Insomnia is prevalent: ~15% of Americans have troubling falling asleep most nights1!
Neuromodulation can help
Notes: Insomnia is a large problem!
Neuromodulation can help!
Graph source: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5353813/

## Slide 4
What is Neuromodulation?
Fall 2023 – Winter 2024
2
  Alteration of nervous signals through application of a stimulus
    Acoustic
    Magnetic
    Transcranial AC, DC
  Goal: Induce correct sleep stage for patient
  Relatively recent, active field of research

## Slide 5
What is Sleep Staging?
Fall 2023 – Winter 2024
3
Inferring sleep stage automatically
Wake
N1
N2
N3
REM
Progressively deeper
Notes: Hypnogram: https://www.ncbi.nlm.nih.gov/books/NBK560713/figure/article-28328.image.f1/

## Slide 6
What is Polysomnography (PSG)?
Fall 2023 – Winter 2024
4
EEG, EOG, EMG, ECG
Heart rate
Airflow
Temperature
Oxygen level
Notes: Polysomnography patient drawing: https://www.psychdb.com/neurology/polysomnography
2:00 min

## Slide 7
What is Polysomnography (PSG)?
Fall 2023 – Winter 2024
4
EEG, EOG, EMG, ECG
Heart rate
Airflow
Temperature
Oxygen level
}
10’s of wires
Cumbersome
Time-consuming
Notes: Polysomnography patient drawing: https://www.psychdb.com/neurology/polysomnography

## Slide 8
Summary of Motivation
Fall 2023 – Winter 2024
5
Insomnia

## Slide 9
Summary of Motivation
Fall 2023 – Winter 2024
5
Insomnia
↓
Neuromodulation

## Slide 10
Summary of Motivation
Fall 2023 – Winter 2024
5
Insomnia
↓
Neuromodulation
↓
Sleep staging

## Slide 11
Summary of Motivation
Fall 2023 – Winter 2024
5
Insomnia
↓
Neuromodulation
↓
Sleep staging
↓
Impractical PSG → Need novel wearable

## Slide 12
Summary of Motivation
Fall 2023 – Winter 2024
5
Insomnia
↓
Neuromodulation
↓
Sleep staging
↓
Impractical PSG → Need novel wearable

## Slide 13
Proposal / Research Question
Fall 2023 – Winter 2024
6
In-ear device performing:
EEG sampling
Automatic sleep staging
Acoustic neuromodulation
Notes: 3:15

## Slide 14
Proposal / Research Question
Fall 2023 – Winter 2024
6
In-ear device performing:
EEG sampling
Automatic sleep staging
Acoustic neuromodulation
Is it feasible?

## Slide 15
Proposal / Research Question
Fall 2023 – Winter 2024
6
In-ear device performing:
EEG sampling
Automatic sleep staging
Acoustic neuromodulation
Develop model + accelerator

## Slide 16
Literature review – Sleep Staging AI
Fall 2023 – Winter 2024
7
Sleep staging using ML since 2017
  CNN, RNN, Transformers
  Most published accuracy > 85%
  Some authors claim problem is “solved”
  Most models have > 1M weights
MultiChannelSleepNet
Notes: MultiChannelSleepNet: https://ieeexplore.ieee.org/stamp/stamp.jsp?tp=&arnumber=10146380
Merges encoder output of multiple EEG channels for higher accuracy
Managed one of the highest accuracy at 87.2%

## Slide 17
Literature review – ASIC Accelerator
Fall 2023 – Winter 2024
8
CPUs: Significant overhead
GPUs: Expensive + power hungry
FPGA: Expensive
ASIC: Most efficient, but inflexible
  First accelerator: DianNao @ 452Gop/s (fixed-point)
DianNao 65nm die
Notes: DianNao: https://users.cs.duke.edu/~lkw34/papers/diannao-asplos2014.pdf

## Slide 18
Literature review – ASIC Accelerator
Fall 2023 – Winter 2024
9
Recent trend: Compute-in-Memory
  Data movement is expensive…
  …put compute near memory!
  CoMeFa: BRAM CiM for 55% reduction in power
Dataflow: Non-von-Neumann architecture → No trad. instructions
CoMeFa
Notes: CoMeFa: https://arxiv.org/pdf/2203.12521.pdf

## Slide 19
Requirements
Fall 2023 – Winter 2024
10
Model
ASIC
Accuracy: ≥ 80%
  Competitive with PSG
Size: < 125kB
  Fit in ASIC
Area: < 2mm2
  Fit in-ear
Effective power: < 5mW
  All-night battery life
Inference latency : < 3s
  Limit phase offset
Notes: CoMeFa: https://arxiv.org/pdf/2203.12521.pdf

## Slide 20
Vision Transformer Model
Fall 2023 – Winter 2024
11
Sequence-to-one
Single encoder
Sleep epoch split in patches
Attention: 8 heads QKV
Window average of 3 softmax

## Slide 21
Accelerator Design
Fall 2023 – Winter 2024
12
Master + 64 CiM architecture
Master
  High-level inference steps
  Interface w/ ADC, mem., RISC-V
Bidirectional bus
  Dense and transpose broadcasts

## Slide 22
Accelerator Design
Fall 2023 – Winter 2024
13
CiM have hardcoded compute
  “Composite” modules share basics
SRAM:
  848 words of intermediate results
  528 words for weights

## Slide 23
Results
Fall 2023 – Winter 2024
14
Model
ASIC
Accuracy: 82.9% (N1/2 merged) ✔
Size: 31589 words ✔
Area: 3.86mm2 ✘
Effective power: 3.05mW ✔
Latency: 6.87ms ✔
fmax: 758MHz

## Slide 24
Results
Fall 2023 – Winter 2024
14
Model
ASIC
Accuracy: 82.9% (N1/2 merged) ✔
Size: 31589 words ✔
Area: 3.86mm2 ✘
Effective power: 3.05mW ✔
Latency: 6.87ms ✔
fmax: 758MHz
ASIC is too large!...

## Slide 25
Analysis and Future Work
Fall 2023 – Winter 2024
15
…but there’s a plan

## Slide 26
Analysis and Future Work
Fall 2023 – Winter 2024
15
…but there’s a plan
Model
Train on STAGES
Additional channels
𝛾 and β elimination

## Slide 27
Analysis and Future Work
Fall 2023 – Winter 2024
15
Galvanic Skin Response
EEG, EMG, EOG, ECG
Blood Volume Pulse
Notes: Patent: https://patentimages.storage.googleapis.com/e2/4d/92/a20ceacf02d9db/US20230225659A1.pdf

## Slide 28
Analysis and Future Work
Fall 2023 – Winter 2024
15
…but there’s a plan
Model
ASIC
Train on STAGES
𝛾 and β elimination
Additional channels
Centralized memory
Centralized compute
Bus elimination
More aggressive power gating
Optimize Q format

## Slide 29
Analysis and Future Work
Fall 2023 – Winter 2024
15
Leakage reduction: - 39.55mW
Area reduction: - 2.11mm2
Latency increase: + 0.305s

## Slide 30
Analysis and Future Work
Fall 2023 – Winter 2024
15
Leakage reduction: - 39.55mW
Area reduction: - 2.11mm2
Latency increase: + 0.305s
}
All requirements met!

## Slide 31
Conclusion
Fall 2023 – Winter 2024
16
Addressing insomnia needs neuromodulation, which needs sleep staging
There is a path towards effective and practical in-ear sleep staging
Results are encouraging…maybe an ASIC isn’t needed
  RISC-V vector extension
  RISC-V + CGRA

## Slide 32
Appendix – Full results
Fall 2023 – Winter 2024
16
Memory results

## Slide 33
Appendix – Full results
Fall 2023 – Winter 2024
16
Compute modules results

## Slide 34
Appendix – Full results
Fall 2023 – Winter 2024
16
System-level results

## Slide 35
Appendix - Acknowledgments
Fall 2023 – Winter 2024
15
Professor Xilin Liu
My father: Claude Robitaille
Computer architecture professors: Mark Jeffrey, Andreas Moshovos, Jason Anderson, Natalie Enright-Jerger and David Halupka
Compute Canada team
Professors Lisa Romkey and Alan Chong for ESC499 organization
Family and friends